Quality management

Final Report

Gauge R&R And PCA    Analysis for E-bike HMI

Professor: Jeh-Nan Pan

Member of group:

Alan Chang(RA7001062)

Bo-wei Chien (R26001100)

Yi-Yang Chuang(R26991096)

Abstract

Process capability indices such as Cp, Cpk, Cpm and Cpmk. The use of process capability indices, we can know that a company can manufacture qualified products ability. But if surveyors or instruments measure ability will be poor, so that process capability that is not credible.

By Gauge R & R experiment to calculate the measurement capability indices about P / T Ratio 、SNR Ratio and DR Ratio, if the measurement system inadequate capability , we must improve the measurement capacity of the surveyors, or replacement of the instruments.

In our experiment, we have 20 pieces each shells and buttons, three surveyors and repeat 2 times.

Our experiments have four comments.

First, the measurement ability is not good because the measure strength of the caliper is different by each measurement of the same operator.

Second, the measurement ability of “Button” is superior to “Shell” because outer diameter is more controllable than inner diameter.

Third, the process capability analysis indicates the production procedure of “Darfon” is excellent.

Forth, the variance of repeatability accounting most to the total Gauge R&R variation

May be due to precision of Instrument is not enough, because our instrument is caliper, it precision only 0.01mm.

But our buttons and shells tolerance are 0.2 mm, so we may be change measure instrument, For example digital micrometer precision is 0.001mm.

We can use micrometer to measuring outer diameter.

Contents

1. Problem Formulation	5

1.1 Motivation	5

1.2 Product introduce	5

1.3 Purpose of Research	5

1.4 Properties of Measurements	6

2. RESEARCH PROCEDURES AND METHOD	7

2.1 Measuring Processes	7

2.2 Cause and effect analysis	8

3. Data collection	9

3.1 Guidelines for the use of the instrument	9

3.2 Products and Instrument of Measurement	9

3.3 Measurement Parameter Determination	10

4. Analysis of results	11

4.1 Descriptive Statistics	11

4.2 R control chart	11

4.3 Residual analysis	13

4.4 ANOVA analysis	15

4.5 Gauge repeatability and reproducibility analysis	16

4.6 Normality test	20

4.7 Process capability analysis	22

4.8 Combination of tolerances	23

4.9 CONCLUSION	23

5. REFERENCES	24

6. APPENDIX	25

Observation of shell	25

Observation of button	26

List of Figures

Figure 1: Flow chart	8

Figure 2: Cause-and-effect diagram	9

Figure 3: Bike HMI	10

Figure 4: Shells and Buttons	10

Figure 5: Specification Tolerance	11

Figure 6: Instrument: Caliper	11

Figure 7: R chart - Shell of first surveyor with outlier	13

Figure 8: R chart - Shell of first surveyor without outlier	13

Figure 9: R chart - Shell of second and third surveyors without outlier	14

Figure 10: R chart - Button of each surveyors without outlier	14

Figure 11: Shell - Normal probability plot of the residuals and the histogram	15

Figure 12: Button - Normal probability plot of the residuals and the histogram	15

Figure 13: Shell - Gauge repeatability and reproducibility analysis.	17

Figure 14: Button - Gauge repeatability and reproducibility analysis.	19

Figure15: Normality test of shell.	22

Figure16: Normality test of Button.	22

List of Tables

Table 1: Descriptive Statistics	12

Table 2: ANOVA of shell	16

Table 3: ANOVA of Button	16

Table 4: Shell’s Gage R&R Study - ANOVA Method	18

Table 5: Shell’s Gage R&R Study	19

Table 6: Button’s Gage R&R Study - ANOVA Method	20

Table 7: Button’s Gage R&R Study	21

Table8: Process capability analysis	23

# 1. Problem Formulation

## 1.1 Motivation

1. The Button of the HMI is the most vulnerable part for product design.

2. The clearance between shell and button will impact the sense of touching for user.

3. The purpose of the experiment is to understand the Gauge R&R and Process Capability Analysis to provide a quality control standard if company has to improve the production process.

## 1.2 Product introduce

The e-bike HMI (Human-Machine-Interface) also called display or console, is the heart for e-bike electrical control system.

The function of the HMI:

1. Battery capacity

2. Information – Time, Speed, Average Speed, Distance & Trip.

3. Power support model – Eco, Speed & Power which one has 3 different power levels.

4. Message – Error code

## 1.3 Purpose of Research

The scope of our research measurement in several directions:

- The precision of measurement be appropriate:

The use of measurement capability index can be seen that the measurement system's ability, which contains the stability of the surveyor measurement capability as well as the measurement instrument of measurement.

- The variation of the product:

Detection of each product’s width is consisted.

- Find not suitable for the measurement member:

R control chart can be seen the measured results, which measured member of the smallest variation, behold the poor.

4. Process capability:

According to the process capability indices can be seen that the manufactured shell and buttons to meet the product specifications

5. Combination of tolerances:

Defective rate of the shell and the button assembly

## 1.4 Properties of Measurements

1. For the same measurement objects and measurement system regardless of present or future should be the same measurement results, that is, should have the "stability"

2. Must be able to produce a close to the true value of the objects to be measured, that is, should have a smaller "bias".

3. The measurement system should be measurable within the accurate measurement results, that is, should have the characteristics of the "linear"

4. Measured by the same surveyor subject to the same parts for repeated measurements, each measurement result should be showing consistent, that is, need to have "repetitive"

5. Measurement of the same subject and location by different trained surveyors, should get the same results of the measurements, that is measurement system should have the "reproducibility"

# 2. RESEARCH PROCEDURES AND METHOD

## 2.1 Measuring Processes

Step.1 - Assign numbers to products and random order surveyor

Step.2 - Create Gauge R&R Study worksheet

Step.3 - Operator 1 2 and 3 measure and record each part measurement

Step.4 - After the first measurement, the same steps for the 2nd

Figure 1: Flow chart

## 2.2 Cause and effect analysis

The general causes of measurement error are man, machine, material, method, measurements and environment. In this experiment, the major measurement error will cause by “measure strength” and “measure angle”.

Figure 2: Cause-and-effect diagram

# 3. Data collection

## 3.1 Guidelines for the use of the instrument

In this experiment, the crew took new HMI shell and buttons in the company, to avoid other factors that affect our experiments, the use of unified experimental guidelines

1 Use the same measuring instruments.

2 After the each measurement was the instrument to zero.

3 Unified measurement locations.

## 3.2 Products and Instrument of Measurement

1. Products: Bike HMI

Figure 3: Bike HMI

2. Measure HMI part: Shell & Button

Figure 4: Shells and Buttons

3. Product Quality Characteristic: Width

4. Specification Tolerance: 0.2 mm

Figure 5: Specification Tolerance

5. Instrument: Caliper

Figure 6: Instrument: Caliper

6. Instrument Precision: 0.01mm

## 3.3 Measurement Parameter Determination

1. Number of Shell & Button: 20 pieces each

2. Number of surveyor: 3

3. Frequency of measurement each surveyor: 2 times

4. Total Frequency of measurement: 3*2*20=120 times of each parts

# 4. Analysis of results

## 4.1 Descriptive Statistics

Twenty samples were used to measure the width and weight of shells and buttons. Three different operators performed the experiment. We first show the descriptive statistics to detect the mean and standard deviation of the width of the shells and buttons, by operator. Then, the total mean and standard deviation were calculated.

Table 1: Descriptive Statistics

## 4.2 R control chart

In statistical quality control, the X and R chart is a type of control chart used to monitor a variable's data when samples are collected at regular intervals from a business or industrial process. When dealing with a quality characteristic that is a continuous variable, it is usually necessary to monitor both the mean value of the quality characteristic and its variability. The process mean can be monitored by using X control chart to ensure the process is in control. The process variability can be monitored by using the with the R control chart. The X and R charts are among the most important and useful on-line statistical process monitoring and control techniques.

When setting up X and R charts, it is better to begin with the R chart, because of the control limits on the X chart will depend on the process variability. Unless the process variability is in control, such limits will not have much meaning. Now, we use R control chart to test whether out-of-control conditions are present. The R charts are shown in Fig.7- Fig .10.

Shell

Figure 7: R chart - Shell of first surveyor with outlier

First Surveyor (one outlier--the ninth one)

The first surveyors R control chart measured shell width of the shell width, one outlier in the ninth Shell sample. So we must remove two measurement values of the first surveyor measuring ninth shell sample and replace by the other two surveyors ninth shell sample mean.

Figure 8: R chart - Shell of first surveyor without outlier

After deleting outlier (9th) and replaced by the second and third surveyor’s 9th shell mean, and then first surveyor’s R –control chart will be in control.

Figure 9: R chart - Shell of second and third surveyors without outlier

We also can see the second and third surveyors R-control charts, the outcome of shell are in control

Button

Figure 10: R chart - Button of each surveyors without outlier

The above shown control charts, also show no presence of an out-of-control condition.

## 4.3 Residual analysis

Before we implement ANOVA analysis, we should check residual is followed some assumptions are equivalent to the assumptions on the response variables (width).

1. The random errors are normally distributed.

2. The random errors have constant variance σ2.

3. The random errors have zero mean.

Shell

Figure 11: Shell - Normal probability plot of the residuals and the histogram

Look at Fig.11, normal probability plot of the residuals and the histogram shows that the residuals follow the normal distribution and scatter randomly. That means that the variance is homogenized. Therefore, the simple linear regression fits this model, and we can use it to predict the shell’s width for different operators and samples.

Button

Figure 12: Button - Normal probability plot of the residuals and the histogram

Look at Fig.12, normal probability plot of the residuals and the histogram shows that the residuals follow the normal distribution and scatter randomly. That means that the variance is homogenized. Therefore, the simple linear regression fits this model, and we can use it to predict the button’s width for different operators and samples.

## 4.4 ANOVA analysis

In this experiment, we have three different operators measuring the width of 15 HMI shell and button twice. The purpose of the experiment is to check whether the different operators, measurement devices and parts will influence the results obtained.

Shell

Table 2: ANOVA of shell

According to the ANOVA Table 2, the width of the shell shows no significant difference between operators (p-value=0.341<0.05), but the part-to-part shows to be significant (p-value=<0.001). And the interaction term between Operator and parts is not significant (p-value=0.399>0.05), so we think that the parts is the factor with more influence on the value of the shell’s width.

Button

Table 3: ANOVA of Button

According to the ANOVA Table 3, the width of the shell shows no significant difference between operators (p-value=0.174<0.05), but the part-to-part shows to be significant (p-value=<0.001). And the interaction term between Operator and parts is not significant (p-value=0.978>0.05), so we think that the parts is the factor with more influence on the value of the shell’s width.

## 4.5 Gauge repeatability and reproducibility analysis

Introduce measurement gauge indicators

P/T ratio

Values of the estimated ratio P/T of 10% or less often are taken to imply adequate gauge capability, estimated ratio P/T between 10%~25% to imply inadequate gauge capability, but it’s acceptable, and estimated ratio P/T of 25% or greater to imply the precision of gauge will be unacceptable.

Signal and noise ratio

AIAG defined the SNR as the number of distinct levels or categories that can be reliably obtained from the measurements. A value of five or greater recommended, and a value of less than two indicates inadequate gauge capability.

Discrimination ratio

Some authors have suggested that for a gauge to be capable the DR must exceed four.

This is a very arbitrary requirement.

Shell

Figure 13: Shell - Gauge repeatability and reproducibility analysis.

(1) Components of Variation

According to Fig.13, the main effect of measurement variability is the Part-to-Part and the Repeatability, which means that the main effect of measurement variability is the auto-deviation of shell and operators.

(2) R chart by Surveyor

The measurement variability of operators is in control.

(3) X bar chart by Surveyor

The measurement values for the three operator’s measurements values are out of control for the X chart. We can also say that the mean of the process is unstable.

(4) Diameter by sample

The diameter shows significant differences between samples.

(5) Diameter by Surveyor

The measurement values for three operators are almost equal, and the interval of measurement value shows no significance difference, that means that there is no significant difference between the three operators.

(6) Surveyor*Sample Interaction

The mean of the measurement shows some difference between samples. We can clearly observe the changes in the width of the shells. The measurements of the three operators follow the same trend. Therefore, there is no interaction between operators and samples.

Table 4: Shell’s Gage R&R Study - ANOVA Method

As following Table.4, this shows variance of repeatability is 17.71%, and variance of reproducibility is 0.07%. Furthermore, variance of Part-To-Part is 82.22%, which is main effect for measurement experiment.

S/N ratio

An estimate of the S/N ratio is 3.04115 small than 5, so precision of gauge will be inadequate.

DR ratio

An estimate of the DR ratio is 10.25575 greater than 4, so precision of gauge will be adequate.

This is a very arbitrary requirement.

Table 5: Shell’s Gage R&R Study

Our number of Distinct Categories is 3.

As following Table.5, this shows stander deviation of repeatability is 42.09%, and stander deviation of reproducibility is 2.62%. Furthermore, variance of Part-To-Part is 90.67%, which is main effect for measurement experiment.

An estimate of the P/T ratio is 26.53%. So precision of gauge will be unacceptable

Button

Figure 14: Button - Gauge repeatability and reproducibility analysis.

(1) Components of Variation

According to Fig.14, the main effect of measurement variability is the Part-to-Part and the Repeatability, which means that the main effect of measurement variability is the auto-deviation of button and operators.

(2) R chart by Surveyor

The measurement variability of operators is in control.

(3) X bar chart by Surveyor

The measurement values for the three operator’s measurements values are out of control for the X chart. We can also say that the mean of the process is unstable.

(4) Diameter by sample

The width shows significant differences between samples.

(5) Diameter by Surveyor

The measurement values for three operators are almost equal, and the interval of measurement value shows no significance difference, that means that there is no significant difference between the three operators.

(6) Surveyor*Sample Interaction

The mean of the measurement shows some difference between samples. We can clearly observe the changes in the width of the buttons. The measurements of the three operators follow the same trend. Therefore, there is no interaction between operators and samples.

Table 6: Button’s Gage R&R Study - ANOVA Method

As following Table 6, this shows variance of repeatability is 13.13%, and variance of reproducibility is 0.07%. Furthermore, variance of Part-To-Part is 86.8%, which is main effect for measurement experiment.

S/N ratio

An estimate of the S/N ratio is 3.626502 small than 5, so precision of gauge will be inadequate.

DR ratio

An estimate of the DR ratio is 14.15152 greater than 4, so precision of gauge will be inadequate.

This is a very arbitrary requirement.

Table 7: Button’s Gage R&R Study

Our number of Distinct Categories is 3.

As following Table 7, this shows stander deviation of repeatability is 36.24%, and stander deviation of reproducibility is 2.59%. Furthermore, variance of Part-To-Part is 93.17%, which is main effect for measurement experiment.

An estimate of the P/T ratio is 20.67% between 10%~25%. So precision of gauge will be inadequate, but it’s acceptable

## 4.6 Normality test

When we use the process capability to analyze the data, we all assume that the data are under normal distribution. Before we analyze the width of shell and button data, we do the normality test. A normal quantile-quantile plot and Shapiro-Wilks test will show that the width of shell and button data is normally distributed.

Shapiro-Wilks test：

H0：the width of shell and button data is normally distributed

H1：the width of shell and button data isn't normally distributed

The data is the width of shell and button, which every operator measure two times. We make normal quantile-quantile plot and Shapiro-Wilks test with different operators' data.

Shell

Figure15: Normality test of shell.

The width of shell data of three operators were calculated p-value of Shapiro-Wilks test, which is greater than 0.1.Under type I error is 0.05, we do not reject H0. The points of data, which drawn on the Q-Q plot (Figure 15), present a straight line. We know the width of shell data is normally distributed, so we can use process capability to analyze the data.

Button

Figure16: Normality test of Button.

The width of button data of three operators were calculated p-value of Shapiro-Wilks test, which is greater than 0.1.Under type I error is 0.05, we do not reject H0. The points of data, which drawn on the Q-Q plot (Figure 16), present a straight line. We know the width of button data is normally distributed, so we can use process capability to analyze the data.

## 4.7 Process capability analysis

The process capability is a measurable property of a process to the specification, expressed as a process capability index.

DARFON Company provided the lower and upper specification limit of width for their HMI product.

So our quality characteristic is the nominal-the-best, then we can calculate Cp,Cpk,Cpm and Cpmk indicates.

Table8: Process capability analysis

The USL & LSL of Shell are 17.35 & 17.15, and the button are 17.05 & 16.85. According to the Cp value, both shell and button are greater than 1.6. Although there is a little bit shift, but the process capability is good, we should keep it up.

## 4.8 Combination of tolerances

In this product, the button will be put into the shell, if the button is much smaller than shell, it will become loosened or disengaged easily. If the shell is much smaller than, it hard to press, even more the button can not be put into the shell. So we calculate the probability of qualified product. The shell and button both follow normal, so the combination of shell and button still follow normal. The probability of X combination greater than 0.2 and smaller than 0.4 is 0.9998476, almost close to 1. The quality of this product is great.

## 4.9 CONCLUSION

- The measurement ability is not good because the measure strength of the caliper is different by each measurement of the same operator.

- The measurement ability of “Button” is superior to “Shell” because outer diameter is more controllable than inner diameter.

- The process capability analysis indicates the production procedure of “Darfon” is excellent.

- The variance of repeatability accounting most to the total Gauge R&R variation.

May be due to precision of Instrument is not enough ,because our instrument is caliper, it precision only 0.01mm.But our buttons and shells tolerance are 0.2 mm, so we may be change measure instrument, For example digital micrometer , it precision is 0.001mm.

We can use micrometer to measuring outer diameter.

# 5. REFERENCES

Douglas C. Montgemery, Introduction to Statistical Quality Control.

潘浙楠、李文瑞，「品質管理」，華泰書局。

江巧玉，「量測系統重複性與再現性的分析研究」。

# 6. APPENDIX

## Observation of shell

| Number | surveyor | observation | Number | surveyor | observation | Number | surveyor | observation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1 | 17.26 | 1 | 2 | 17.25 | 1 | 3 | 17.25 |
| 1 | 1 | 17.27 | 1 | 2 | 17.26 | 1 | 3 | 17.24 |
| 2 | 1 | 17.26 | 2 | 2 | 17.24 | 2 | 3 | 17.24 |
| 2 | 1 | 17.24 | 2 | 2 | 17.25 | 2 | 3 | 17.25 |
| 3 | 1 | 17.30 | 3 | 2 | 17.29 | 3 | 3 | 17.30 |
| 3 | 1 | 17.30 | 3 | 2 | 17.30 | 3 | 3 | 17.31 |
| 4 | 1 | 17.29 | 4 | 2 | 17.29 | 4 | 3 | 17.29 |
| 4 | 1 | 17.27 | 4 | 2 | 17.27 | 4 | 3 | 17.29 |
| 5 | 1 | 17.27 | 5 | 2 | 17.26 | 5 | 3 | 17.26 |
| 5 | 1 | 17.26 | 5 | 2 | 17.26 | 5 | 3 | 17.27 |
| 6 | 1 | 17.22 | 6 | 2 | 17.22 | 6 | 3 | 17.20 |
| 6 | 1 | 17.21 | 6 | 2 | 17.23 | 6 | 3 | 17.22 |
| 7 | 1 | 17.25 | 7 | 2 | 17.25 | 7 | 3 | 17.27 |
| 7 | 1 | 17.24 | 7 | 2 | 17.26 | 7 | 3 | 17.25 |
| 8 | 1 | 17.25 | 8 | 2 | 17.24 | 8 | 3 | 17.25 |
| 8 | 1 | 17.23 | 8 | 2 | 17.26 | 8 | 3 | 17.26 |
| 9 | 1 | 17.23 | 9 | 2 | 17.23 | 9 | 3 | 17.23 |
| 9 | 1 | 17.23 | 9 | 2 | 17.24 | 9 | 3 | 17.22 |
| 10 | 1 | 17.26 | 10 | 2 | 17.25 | 10 | 3 | 17.26 |
| 10 | 1 | 17.25 | 10 | 2 | 17.24 | 10 | 3 | 17.26 |
| 11 | 1 | 17.28 | 11 | 2 | 17.26 | 11 | 3 | 17.27 |
| 11 | 1 | 17.27 | 11 | 2 | 17.27 | 11 | 3 | 17.25 |
| 12 | 1 | 17.25 | 12 | 2 | 17.24 | 12 | 3 | 17.24 |
| 12 | 1 | 17.25 | 12 | 2 | 17.24 | 12 | 3 | 17.23 |
| 13 | 1 | 17.26 | 13 | 2 | 17.24 | 13 | 3 | 17.24 |
| 13 | 1 | 17.26 | 13 | 2 | 17.25 | 13 | 3 | 17.25 |
| 14 | 1 | 17.23 | 14 | 2 | 17.25 | 14 | 3 | 17.25 |
| 14 | 1 | 17.24 | 14 | 2 | 17.25 | 14 | 3 | 17.22 |
| 15 | 1 | 17.26 | 15 | 2 | 17.25 | 15 | 3 | 17.26 |
| 15 | 1 | 17.25 | 15 | 2 | 17.27 | 15 | 3 | 17.25 |
| 16 | 1 | 17.27 | 16 | 2 | 17.26 | 16 | 3 | 17.26 |
| 16 | 1 | 17.26 | 16 | 2 | 17.26 | 16 | 3 | 17.26 |
| 17 | 1 | 17.25 | 17 | 2 | 17.27 | 17 | 3 | 17.26 |
| 17 | 1 | 17.26 | 17 | 2 | 17.25 | 17 | 3 | 17.27 |
| 18 | 1 | 17.29 | 18 | 2 | 17.29 | 18 | 3 | 17.28 |
| Number | surveyor | observation | Number | surveyor | observation | Number | surveyor | observation |
| 18 | 1 | 17.30 | 18 | 2 | 17.28 | 18 | 3 | 17.30 |
| 19 | 1 | 17.26 | 19 | 2 | 17.24 | 19 | 3 | 17.25 |
| 19 | 1 | 17.26 | 19 | 2 | 17.25 | 19 | 3 | 17.26 |
| 20 | 1 | 17.24 | 20 | 2 | 17.24 | 20 | 3 | 17.25 |
| 20 | 1 | 17.26 | 20 | 2 | 17.25 | 20 | 3 | 17.26 |

## Observation of button

| Number | surveyor | observation | Number | surveyor | observation | Number | surveyor | observation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1 | 16.96 | 1 | 2 | 16.95 | 1 | 3 | 16.96 |
| 1 | 1 | 16.99 | 1 | 2 | 16.96 | 1 | 3 | 16.95 |
| 2 | 1 | 16.96 | 2 | 2 | 16.97 | 2 | 3 | 16.95 |
| 2 | 1 | 16.95 | 2 | 2 | 16.96 | 2 | 3 | 16.93 |
| 3 | 1 | 16.95 | 3 | 2 | 16.95 | 3 | 3 | 16.95 |
| 3 | 1 | 16.92 | 3 | 2 | 16.95 | 3 | 3 | 16.98 |
| 4 | 1 | 16.965 | 4 | 2 | 16.97 | 4 | 3 | 16.96 |
| 4 | 1 | 16.95 | 4 | 2 | 16.98 | 4 | 3 | 16.91 |
| 5 | 1 | 16.93 | 5 | 2 | 16.97 | 5 | 3 | 16.96 |
| 5 | 1 | 16.93 | 5 | 2 | 16.95 | 5 | 3 | 16.96 |
| 6 | 1 | 16.94 | 6 | 2 | 16.97 | 6 | 3 | 16.94 |
| 6 | 1 | 16.98 | 6 | 2 | 16.92 | 6 | 3 | 16.97 |
| 7 | 1 | 16.96 | 7 | 2 | 16.99 | 7 | 3 | 16.95 |
| 7 | 1 | 16.92 | 7 | 2 | 16.96 | 7 | 3 | 16.95 |
| 8 | 1 | 16.94 | 8 | 2 | 16.92 | 8 | 3 | 16.96 |
| 8 | 1 | 16.94 | 8 | 2 | 16.94 | 8 | 3 | 16.93 |
| 9 | 1 | 16.95 | 9 | 2 | 16.9575 | 9 | 3 | 16.92 |
| 9 | 1 | 16.96 | 9 | 2 | 16.92 | 9 | 3 | 16.95 |
| 10 | 1 | 16.95 | 10 | 2 | 16.91 | 10 | 3 | 16.965 |
| 10 | 1 | 16.95 | 10 | 2 | 16.96 | 10 | 3 | 16.95 |
| 11 | 1 | 16.93 | 11 | 2 | 16.97 | 11 | 3 | 16.95 |
| 11 | 1 | 16.95 | 11 | 2 | 16.95 | 11 | 3 | 16.93 |
| 12 | 1 | 16.98 | 12 | 2 | 16.93 | 12 | 3 | 16.96 |
| 12 | 1 | 16.96 | 12 | 2 | 16.97 | 12 | 3 | 16.96 |
| 13 | 1 | 16.91 | 13 | 2 | 16.99 | 13 | 3 | 16.94 |
| 13 | 1 | 16.96 | 13 | 2 | 16.95 | 13 | 3 | 16.98 |
| 14 | 1 | 16.96 | 14 | 2 | 16.96 | 14 | 3 | 16.98 |
| 14 | 1 | 16.94 | 14 | 2 | 16.94 | 14 | 3 | 16.96 |
| Number | surveyor | observation | Number | surveyor | observation | Number | surveyor | observation |
| 15 | 1 | 16.97 | 15 | 2 | 16.94 | 15 | 3 | 16.95 |
| 15 | 1 | 16.95 | 15 | 2 | 16.91 | 15 | 3 | 16.93 |
| 16 | 1 | 16.95 | 16 | 2 | 16.95 | 16 | 3 | 16.96 |
| 16 | 1 | 16.96 | 16 | 2 | 16.96 | 16 | 3 | 16.92 |
| 17 | 1 | 16.93 | 17 | 2 | 16.94 | 17 | 3 | 16.9575 |
| 17 | 1 | 16.92 | 17 | 2 | 16.93 | 17 | 3 | 16.98 |
| 18 | 1 | 16.95 | 18 | 2 | 16.95 | 18 | 3 | 16.94 |
| 18 | 1 | 16.965 | 18 | 2 | 16.92 | 18 | 3 | 16.95 |
| 19 | 1 | 16.95 | 19 | 2 | 16.98 | 19 | 3 | 16.93 |
| 19 | 1 | 16.96 | 19 | 2 | 16.95 | 19 | 3 | 16.95 |
| 20 | 1 | 16.99 | 20 | 2 | 16.95 | 20 | 3 | 16.95 |
| 20 | 1 | 16.96 | 20 | 2 | 16.96 | 20 | 3 | 16.94 |